## Cover Sheet
PAPR Design Review
June 8, 2020

## NIOSH DESIGN REQUIREMENTS
NIOSH Requirements per 42 CFR Part 84
84.61 | General Construction Requirements | Met Yes/No | Notes
b) | Respirator components which come into contact with the wearer's skin shall be made of nonirritating materials. | yes | Hood made of Tychem. Currently used in hospitals. `
d) | Mouthpieces, hoods, helmets, and facepieces, except those employed in single-use respirators, shall be constructed of materials which will withstand repeated disinfection as recommended by the applicant in his instructions for use of the device. | Yes | Hood good for 100 cycles
84.62 | Component parts; minimum requirements
a) | The component parts of each respirator shall be:
1 | Designed, constructed, and fitted to insure against creation of any hazard to the wearer; | Yes | Hood made of Tychem. Currently used in hospitals. Non abrasive material.
2 | Assembled to permit easy access for inspection and repair of functional parts; and |  | Each component of the unit can be disassembled for inspection.
3 | Assembled to permit easy access to parts which require periodic cleaning and disinfecting. |  | Each component of the unit can be disassembled for inspection.
b) | Replacement parts shall be designed and constructed to permit easy installation and to maintain the effectiveness of the respirator. |  | Battery and hood can be considered replaceable parts. Battery is stored in an external pouch and can be swapped as frequently as needed by disconnecting from the connector. Hood can be replaced in the same manner.
84.171 | Required components and attributes
The components of each air-purifying particulate respirator must meet the minimum construction requirements set forth in subpart G of this part. Each air-purifying particulate respirator described in 84.170 must, where its design requires, contain the following component parts:
a) | Respiratory inlet covering |  | Inlet covering is the hood
1) | Tight-fitting respiratory inlet coverings must be designed and constructed to fit persons with various facial shapes and sizes either: | N/A | Loose fitting hood
i) | By providing more than one size: or | N/A
ii) | By providing more than one size which will fit varying facial shapes and sizes. | N/A
2) | Full facepieces must provide for optional use of corrective spectacles or lenses, which must not reduce the respiratory protective qualities of the respirator. | N/A
3) | Loose fitting respiratory inlet coverings must be designed and constructed to fit persons with various head sizes, provide for the optional use of corrective spectacles or lenses, and insure against any restriction of movement by the wearer. |  | Adult head-sized hood. The hood may be used with or without corrective lenses. Adequate space inside for glasses to fit.
5) | Respiratory inlet coverings that incorporate a lens or face shield must be designed to prevent eyepiece fogging. |  | There is positive airflow through the hood. There is no buildup of air inside the hood. Exhalation air is pushed down and out of the hood from the wearer's nose and should minimize moisture buildup inside the hood. This can be physically tested this by wearing it.
7) | Respiratory inlet coverings must be designed and constructed to provide adequate vision which is not distorted by the eyepieces. |  | Uses clear plastic lenses. Tested by Mike and Aaron and no distortions were noted.
b) | Filter unit. The respirator manufacturer, as part of the application for certification, must specify the filter series and the filter efficiency level (i.e., “N95,” “R95,” “P95,” “N99,” “R99,” “P99,” “N100,” “R100,” “P100,” “HE,” “PAPR100-N” or “PAPR100-P”) for which certification is being sought. |  | PAPR100-N
2) | Filters for powered respirators (classes HE and PAPR100) must be prominently labeled as follows:
ii) | PAPR100-N filters must be labeled as “PAPR100-N Particulate Filter (99.97% filter efficiency level)” and must be color coded magenta. |  | Label to be created. Will be magenta and correctly marked.
iii) | PAPR100-P filters must be labeled as “PAPR100-P Particulate Filter (99.97% filter efficiency level)” and must be color coded magenta. |  | N/A
c) | Valves.
1) | Inhalation and exhalation valves must be protected against distortion. | N/A | No valve
2) | Inhalation valves must be designed and constructed and provided where necessary to prevent excessive exhaled air from adversely affecting filters, except where filters are specifically designed to resist moisture. | N/A | Exhaled air is forced downwards and out of hood by positive air flow coming from top of hood.
3) | Exhalation valves must be: | N/A
i) | Provided where necessary; | N/A
ii) | Protected against damage and external influence; and | N/A
iii) | Designed and constructed to prevent inward leakage of contaminated air. | N/A
d) | Head harness.
1) | All facepieces must be equipped with head harnesses designed and constructed to provide adequate tension during use and an even distribution of pressure over the entire area in contact with the face. | N/A
2) | Facepiece head harnesses, except those employed on filtering facepiece respirators, must be adjustable and replaceable. | N/A
e) | Breathing tube. Flexible breathing tubes used in conjunction with respirators must be designed and constructed to prevent: |  | Spring coiled design prevents kinking and collapse of hose.
1) | Restriction of free head movement; |  | Hood freely rotates between the hood the hose connection.
2) | Disturbance of the fit of facepieces, mouthpieces, or loose fitting respiratory-inlet covering; |  | No facepiece or mouth piece. The hose attaches to the back of the hood and allows for free rotating around the hose connection.
3) | Interference with the wearer's activities; and |  | Hose comes out of the PAPR at an angle to follow the streamline of the wearer's back. The hose does not stick out and will not interfere with the wearer's activities.
4) | Shutoff of airflow due to kinking, or from chin or arm pressure. |  | Anti-kink hose. Low flow warning present. If flow rate drops in the hood, the fan speed will increase to compensate. If the unit goes into a catastrophic uncontrollable state, warning will sound and unit will turn off.
g) | Container. |  | This is the container the unit will be transported in.
1) | Except as provided in paragraph (b) of this section, each respirator must be equipped with a substantial, durable container bearing markings which show the applicant's name, the type of respirator it contains, and all appropriate approval labels. |  | Will be a hard case, to be determined.
2) | Containers for respirators may provide for storage of more than one respirator; however, such containers must be designed and constructed to prevent contamination of respirators which are not removed, and to prevent damage to respirators during transit. |  | N/A, Containers will only be for one PAPR unit per time
h) | Harness.
1) | Each respirator must, where necessary, be equipped with a suitable harness designed and constructed to hold the components of the respirator in position against the wearer's body. |  | There is a belt with a buckle. Adjustable to make the belt tighter.
2) | Harnesses must be designed and constructed to permit easy removal and replacement of respirator parts, and, where applicable, provide for holding a full facepiece in the ready position when not in use. |  | Belt has buckle and can be easily disassembled. We do not use a full facepiece.
i) | Attached blower—PAPR classes HE and PAPR100. Blowers must be designed to achieve the air flow rates required by the testing standards in §84.175. |  | Required flowrate > 6 CFM. Prototype has already been tested at 16 CFM.
j) | Low-flow warning device—PAPR class PAPR100.
1) | The design must include a low-flow warning. It must actively and readily indicate when flow inside the respiratory inlet covering falls below the minimum air flow defined in §84.175. |  | In design, but intent is to use pressure transducers to measure flow rate. A warning will be audible via speakers when flowrate drops below 4 CFM.
2) | Any warning must be detectable by the wearer without any intervention by the wearer. |  | Warning sound to be sufficiently above 80 decibels.
3) | Warning devices must be configured so that they may not be de-energized while the blower is energized. |  | Hardware will take care of this. Power will run parallel to the warnings and will not be de-energized when the fan is running. Warning is dependent on low battery signal or signal from flow measurements.
4) | During use, warning devices must not switch off automatically and must not be capable of being switched off by the wearer. |  | Software controlled, cannot be tampered by the user. Buttons need to be pressed down for a certain amount of time to input change to prevent accidental presses.
5) | Any warnings which require different reactions by the wearer must be distinguishable from one another. |  | Different sound for low battery vs flow rate/failure beep
6) | If the warning provided is audible only, or other warnings are not readily apparent to the wearer, the minimum sound level must be 80 dBA. |  | Sound intended to be above 80 decibels.
Regulatory Design Requirements for Battery-Powered PAPR for use in Healthcare Environments NIOSH Requirements per 42 CFR Part 84

## Ref 84.170
Reference
84.17
a) | Non-powered air-purifying particulate respirators (series N, R, and P).
1) | Non-powered air-purifying particulate respirators utilize the wearer's negative inhalation pressure to draw the ambient air through the air-purifying filter elements (filters) to remove particulates from the ambient air. They are designed for use as respiratory protection against atmospheres with particulate contaminants at concentrations that are not immediately dangerous to life or health and that contain adequate oxygen to support life.
2) | Non-powered air-purifying particulate respirators are classified into three series, N-, R-, and P-series. The N-series filters are restricted to use in those workplaces free of oil aerosols. The R- and P-series filters are intended for removal of any particulate that includes oil-based liquid particulates.
3) | Non-powered air-purifying particulate respirators are classified according to the efficiency level of the filter(s) as tested according to the requirements of this part.
i) | N100, R100, and P100 filters must demonstrate a minimum efficiency level of 99.97 percent.
ii) | N99, R99, and P99 filters must demonstrate a minimum efficiency level of 99 percent.
iii) | N95, R95, and P95 filters must demonstrate a minimum efficiency level of 95 percent.
b) | Powered air-purifying particulate respirators (PAPR classes HE and PAPR100).
1) | Powered air-purifying particulate respirators utilize a blower to move the ambient air through the air-purifying filter elements (filters) to remove particulate contaminants and deliver clean air to the respiratory inlet covering. They are designed for use as respiratory protection against atmospheres considered not immediately dangerous to life or health and that contain adequate oxygen to support life.
2) | Powered air-purifying particulate respirators are classified into two classes, HE and PAPR100, and three series, HE, PAPR100-N, and PAPR100-P. The N-series filters are restricted to use in those workplaces free of oil aerosols. The P-series filters are intended for removal of any particulate that includes oil-based liquid particulates.
3) | All three filter series, HE, PAPR100-N, and PAPR100-P, for powered air-purifying particulate respirators must demonstrate a minimum efficiency level of 99.97 percent.

## Subpart G
Subpart G
General Performance and Construction Requirements
§84.60 Construction and performance requirements; general.
(a) The Institute shall issue approvals for the types of respirators described in subparts H through O of this part which have met the minimum requirements set forth for such respirators in this part.
(b) In addition to the types of respirators specified in subparts H through L of this part, the Institute shall issue approvals for other respiratory protective devices not specifically described in this part subject to such additional requirements as may be imposed in accordance with §84.63(c).
[60 FR 30355, June 8, 1995, as amended at 77 FR 14192, Mar. 8, 2012; 85 FR 20607, Apr. 14, 2020]
§84.61 General construction requirements.
(a) Respirators will not be accepted by the Institute for examination, inspection and testing unless they are designed on sound engineering and scientific principles, constructed of suitable materials and evidence good workmanship.
(b) Respirator components which come into contact with the wearer's skin shall be made of nonirritating materials.
(c) Components replaced during or after use shall be constructed of materials which will not be damaged by normal handling.
(d) Mouthpieces, hoods, helmets, and facepieces, except those employed in single-use respirators, shall be constructed of materials which will withstand repeated disinfection as recommended by the applicant in his instructions for use of the device.
§84.62 Component parts; minimum requirements.
(a) The component parts of each respirator shall be:
(1) Designed, constructed, and fitted to insure against creation of any hazard to the wearer;
(2) Assembled to permit easy access for inspection and repair of functional parts; and
(3) Assembled to permit easy access to parts which require periodic cleaning and disinfecting.
(b) Replacement parts shall be designed and constructed to permit easy installation and to maintain the effectiveness of the respirator.
§84.63 Test requirements; general.
a) Each respirator and respirator component shall when tested by the applicant and by the Institute, and meet the applicable requirements set forth in subparts H through O of this part.
(b) Where a combination respirator is assembled from two or more types of respirators, as described in this part, each of the individual respirator types which have been combined shall, as applicable, meet the minimum requirements for such respirators set forth in subparts H through O of this part, and such combination respirators, except as specified in §84.70(b)(2), will be classified by the type of respirator in the combination which provides the least protection to the user.
(c) In addition to the minimum requirements set forth in subparts H through O of this part, the Institute reserves the right to require, as a further condition of approval, any additional requirements deemed necessary to establish the quality, effectiveness, and safety of any respirator used as protection against hazardous atmospheres.
(d) Where it is determined after receipt of an application that additional requirements will be required for approval, the Institute will notify the applicant in writing of these additional requirements, and necessary examinations, inspections, or tests, stating generally the reasons for such requirements, examinations, inspections, or tests.
[60 FR 30355, June 8, 1995, as amended at 77 FR 14192, Mar. 8, 2012; 85 FR 20607, Apr. 14, 2020]
§84.64 Pretesting by applicant; approval of test methods.
(a) Prior to making or filing any application for approval or modification of approval, the applicant shall conduct, or cause to be conducted, examinations, inspections, and tests of respirator performance which are equal to or exceed the severity of those prescribed in this part.
(b) With the application, the applicant shall provide a statement to the Institute showing the types and results of the examinations, inspections, and tests required under paragraph (a) of this section and state that the respirator meets the minimum requirements of subparts H through O of this part, as applicable. Complete examination, inspection, and test data shall be retained on file by the applicant and be submitted, upon request, to the Institute.
(c) The Institute may, upon written request by the applicant, provide drawings and descriptions of its test equipment and otherwise assist the applicant in establishing a test laboratory or securing the services of a testing agency.
(d) No approval will be issued until the Institute has validated the applicant's test results.
[60 FR 30355, June 8, 1995, as amended at 77 FR 14193, Mar. 8, 2012; 85 FR 20607, Apr. 14, 2020]
§84.65 Conduct of examinations, inspections, and tests by the Institute; assistance by applicant; observers; recorded data; public demonstrations.
(a) All examinations, inspections, and tests conducted pursuant to subparts H through O of this part will be under the sole direction and control of the Institute.
(b) The Institute may, as a condition of approval, require the assistance of the applicant or agents of the applicant during the assembly, disassembly, or preparation of any respirator or respirator component prior to testing or in the operation of such equipment during testing.
(c) Only Institute personnel, persons assisting the Institute pursuant to paragraph (b) of this section, and such other persons as are requested by the Institute or the applicant to be observers, shall be present during any examination, inspection, or test conducted prior to the issuance of an approval by the Institute for the equipment under consideration.
(d) The Institute shall hold as confidential any analyses, drawings, specifications, or materials submitted by the applicant and shall not disclose any principles or patentable features of such equipment, except as required by statute or regulation.
(e) As a condition of each approval issued for any respirator, the Institute reserves the right, following the issuance of such approval, to conduct such public tests and demonstrations of the approved respiratory equipment as is deemed appropriate.
[60 FR 30355, June 8, 1995, as amended at 77 FR 14193, Mar. 8, 2012; 85 FR 20607, Apr. 14, 2020]
§84.66 Withdrawal of applications.
(a) Any applicant may, upon a written request submitted to the Institute, withdraw any application for approval of any respirator.
(b) Upon the receipt of a written request from the applicant for the withdrawal of an application, NIOSH will invoice the applicant based on the fee calculated, as specified under §84.21(e).
[60 FR 30355, June 8, 1995, as amended at 80 FR 3907, Jan. 26, 2015]