## Sheet1
|  | Equipment Required: | * Leak tester/ Particle Emitter * Torque reader * Decibel Meter/equiv. testing equipment * Stress test/Torque pull tester/Shear tester * PCB test equipment * TSI machine/ Filter test equipment * Airflow test equipment (including reader) * Weigh scale?
Item: | Description: | Requesting?: | Inspection: | Test Method Preferred Option: | Test Method Backup to Preferred Option: | Classification: | AQLs: | Production Phase: | Location: | AQL: | Recorded Results Area:
1 | BATTERY CHARGER * No longer a charger * Is a cord * AND power source? | YES | Assembly is Correct, Clean & in Good Condition (All parts are present per spec and free of: rusting, leaking, oils, residue, discoloration, dents, cracks, chips, holes, flaking, etc.) | Visual Inspection | N/A | Minor | 4 | Pre-Production
|  | YES | Functioning (All parts are functional when checked with ____. Charger assembly activates the charging of the PAPR battery when connected to the airbox.) | Electronic Tests | N/A | Minor | 4 | Any
2 | BATTERY * Specify for Battery PACK now | YES | Assembly is Correct, Clean & in Good Condition (All parts are present per spec and free of: rusting, leaking, oils, residue, discoloration, dents, cracks, chips, holes, flaking, etc.) | Visual Inspection | N/A | Minor | 4 | Pre-Production
|  | YES | Functioning (Battery/batteries are functional when tested and are charged to specs within tolerances.) | Electronic Tests | N/A | Major B | 2.5 | Any
3 | LABEL & IFU PRODUCTION | YES | Instructions For Use (IFU): Assembly is Correct (Contains correct information per specs, are of correct size per specs.) | Visual Inspection | Weight Inspection? Other if recommended | Minor | 4 | In-Production
|  | YES | Labels in Good Condition (Labels are readable and printed per specs (correct material, ink quality, color, non-smudging properties), and free of: print issues (ex. barring, blotches), smudges, and other damage (ex. tears, "dog-ears". This remains so after labels are applied to their associated components.) | Visual Inspection | Other if recommended | Minor | 4 | In-Production OR Post-Production
|  | YES | Labels: Assembly is Correct Part 1 (Labels contain the correct information per specs, and are of correct size and color per specs.) | Visual Inspection | Other if recommended | Minor | 4 | In-Production OR Post-Production
|  | YES | Labels: Assembly is Correct Part 2 (The required labels (including Full Approval Label, Abbreviated Approval Label, Container Label) are present in their required locations (per specs), straight, and are attached securely onto smooth, undisturbed surfaces. There is no evidence of potential for damage or removal of the labels once they are in position (labels do not slide, no bubbling, no lifted parts, nothing underneath which prevents smooth contact between the label and the surface to which it is adhered to).) | Visual Inspection | Other if recommended | Minor | 4 | In-Production OR Post-Production
|  | YES | Instructions For Use (IFU): Assembly is Correct (Contains correct number of pages, information is printed per specs, free of blanks where information is supposed to be present.) | Visual Inspection | Weight Inspection? Other if recommended | Minor | 4 | In-Production OR Post-Production
|  | YES | Instructions For Use (IFU): Product is in Good Condition (Binding is secure & pages are intact, pages are readable and printed per specs (paper, ink quality, color), free of: smudges, print issues (ex. Barring, etc.), tears, "dog-ears," etc.) , is straight, adheres securely without sliding or peeling, & is free of: lifted edges, air bubbles, smudges.) | Visual Inspection | Other if recommended | Minor | 4 | In-Production OR Post-Production
4 | ACCESSORIES | YES | Assembly is Correct (All accessories are present per spec.) | Visual Inspection | N/A | Minor | 4 | Post-Production
|  | YES | Cleanliness & Condition (All accessories are free of: debris, cracks, holes, flaking, chips, dents, stretching, stretches, etc.) | Visual Inspection | N/A | Minor | 4 | Pre-Production
5 | HOSE | YES | Assembly is Correct (Part 1) (Hose is present, attached to airbox, and has hood connector at opposite end. All, parts securely attached.) | Visual Inspection: parts present | N/A | Minor | 4 | Post-Production OR In-Production
|  | YES | Assembly is Correct (Part 2) (Hose is present, attached to airbox, and has hood connector at opposite end. All, parts securely attached.) | As part of "Full Assembly" Leak test | Leak test/Particle Emitter test | Major A | 1 | Post-Production
|  | YES | Cleanliness & Condition (Part 1) (Free of cracks, holes, flaking, chips, dents, stretching / kinks.) | Visual Inspection | Supplier's CoC | Major A | 1 | Pre-Production
|  | YES | Cleanliness & Condition (Part 2) (Supplier CoC shows the hose has been made per spec) | Supplier CoC | N/A | Minor | 4 | Pre-Production
6 | FILTER | YES | Assembly is Correct (Part 1) (Filter gasket sits securely within the airbox & forms a valid seal on all sides: no gaps are present, there is a flush seal, & filter sits flat against plenum/manifold.) | Visual Inspection | N/A | Major A | 1 | Post-Production
|  | YES | Assembly is Correct (Part 2) (Filter gasket sits securely within the airbox & forms a valid seal on all sides: no gaps are present, there is a flush seal, & filter sits flat against plenum/manifold.) | Leak Test (reverse airflow through hose?) | N/A | Major A | 1 | Post-Production
|  | YES | Filter Efficiency "Determination of Particulate Filter Penetration (PAPR) Test." (sampling level verification available?) | TEB-APR-STP-0001 (or equivalent) | Supplier's CoC (but best if Manufacturer can do test if Supplier does not verify with NIOSH-approved test methods) >> Potential to have Supplier agreement that filters maintain NIOSH approval? | Major A | 1 | Pre-Production
|  | YES | Silica Dust Loading "Determination of Silica Dust Loading for PAPR Filters Test." | RCT-APR-STP-0025 (or equivalent) | Supplier's CoC (but best if Manufacturer can do test if Supplier does not verify with NIOSH-approved test methods) >> Potential to have Supplier agreement that filters maintain NIOSH approval? | Major A | 1 | Pre-Production
|  | YES (shows if unpacking & handling issue) (ensures integrity maintained on all sides of filter pre-installation/post -unpackaging and -handling) | Cleanliness & Condition (Part 1) (directly before installation) (Filter media must be free of tears, holes, flakes, dust, debris & sealed all the way to edge of its gasket.) | Visual Inspection | Supplier's CoC (Warning: does not ensure against probability of damage during the unpackaging and handling stage at the Manufacturing site. The delicate nature of filter media must be strongly considered in determining appropriate inspection points, as should the filter's importance as the foundational source of protection in the PAPR) | Major A | 1 | In-Production (directly before installation)
|  | YES (shows if installation issue) (double checks integrity is maintained -as visible on front face of filter- during installation) | Cleanliness & Condition (Part 2) (directly before case face is fastened to rear case) (Filter media must be free of tears, holes, flakes, dust, debris & sealed all the way to edge of its gasket.) | Visual Inspection | Supplier's CoC (Warning: does not ensure against probability of damage during the unpackaging and handling stage at the Manufacturing site. The delicate nature of filter media must be strongly considered in determining appropriate inspection points, as should the filter's importance as the foundational source of protection in the PAPR) | Major A | 1 | In-Production (directly before case face is fastened to rear case)
|  | IDEALLY | Cleanliness & Condition (Part 3) (after receiving and unpackaging) (Filter media must be free of tears, holes, flakes, dust, debris & sealed all the way to edge of its gasket.) | Visual Inspection | Supplier's CoC (Warning: does not ensure against probability of damage during the packaging & shipping stages at the Supplier site, nor during the unpackaging stage at the Manufacturing site. The delicate nature of filter media must be strongly considered in determining appropriate inspection points, as should the filter's importance as the foundational source of protection in the PAPR) | Major A | 1 | Pre-Production
7 | FAN | YES | Assembly Correct (Part 1): Fan placement (Fan is placed in correct orientation.) | Visual Inspection | Post-Production flow/function test of full assembly (Warning: increases costs if issues found at this point vs. earlier) | Major A | 1 | In-Production
|  | UNSURE | Assembly Correct (Part 2): Wiring (Wiring for fan is correct.) | Electronic Tests | Visual Inspection | Major A | 1 | In-Production
|  | YES | Assembly Correct (Part 3) Gaskets & Backing Plate ( Fan gasket, sleeve, & backing plate are in place, correctly seated, & form a proper seal.) | Leak test/Particle Emitter test | Visual Inspection (Warning: does not verify integrity of seal against leakage from rear-case air entering filtered air path) | Major A | 1 | In-Production
|  | YES | Assembly Correct (Part 4) Fasteners Secure (Fan/blower mounting screws are threaded properly into backing plate (plate inside plenum to receive screws), and tightened at appropriate/specified torque level) | Torque Limiter / Torque Reader test | N/A | Major A | 1 | In-Production
|  | YES | Assembly Correct (Part 5) Fasteners Sealed (Screws which require seals (per spec's) are sealed and sealed securely.) | Leak test/Particle Emitter test | N/A | Major A | 1 | In-Production
|  | YES | Cleanliness & Condition (Fan and its backing plate, sleeves, adapters & gaskets are free of cracks, tears, holes, flaking, chipping, debris, or other damage which might compromise integrity of the parts.) | Visual Inspection | Supplier CoC (Warning: does not ensure against damage and debris which could compromise the device's long-term protection capacity) | Major A | 1 | Pre-Production OR In-Production
|  | YES | Airflow Minimum Maintained "Determination of Air Flow PAPR Test." | RCT-APR-STP-0012 (or equivalent) | N/A | Major A | 1 | Post-Production
|  | NO (Not needed for Loose-Fitting Facepiece) | Noise Test "Determination of Noise Level Test, Powered Air-Purifying Respirator With Hoods or Helmets." | RCT-APR-STP-0030 (or equivalent) | N/A | Major B | 2.5 | Post-Production (with hood)
8 | FULL ASSEMBLY | YES | Exterior is Leak-Proof (External unit is fully sealed/leak-free, except to allow air through the air inlet and through the designated airpath through the hose to the hood. Includes: Full case exterior (Case, Case Closure Seam, UI Membrane, Case-to-UI Membrane Ridge, O-rings?, Exterior Screws), Case-to-Hose Coupling, Hose, Hose-to-Hood Coupling.) | Leak test/Particle Emitter test | N/A | Major A | 1 | Post-Production
9 | CASE / ENCLOSURE | YES | Buttons are functional with membrane attached (All lights turns on and off when pressed for each action.) | Visual Inspection | N/A | Major B | 2.5 | Post-Production
|  | UNSURE | Gaskets: Cleanliness & Condition (Gasket are free of cracks, tears, holes, flaking, chipping, debris, or other damage which might compromise integrity of the parts.) | Visual Inspection | Supplier CoC (Warning: does not ensure against damage and debris which could compromise the device's long-term protection capacity) | Major A | 1 | Pre-Production OR In-Production
|  | YES | Plenum Sealed (Plenum is properly sealed to adjacent casing.) | Leak Test (reverse airflow through hose?) | Visual Inspection (Warning: does not verify integrity of seal against leakage from rear-case air entering filtered air path) | Major A | 1 | In-Production
|  | YES (checks size specs & checks if closure issue is due to case fabrication vs. filter installation) | Empty Case Closes (Case closes easily but securely without the filter.) | Visual Inspection | N/A | Minor | 4 | In-Production
|  | YES | Assembly Correct (Part 1) All Parts Connected (Case is complete with all pieces/parts per specs.) | Visual Inspection | Weight Inspection | Major A | 1 | Post-Production
|  | YES | Assembly Correct (Part 2 ) Main Case Fasteners Secure (Main case assembling screws: properly threaded into mating threads and torqued to spec's.) | Torque Limiter/Torque Reader test | N/A | Major A | 1 | In-Production OR Post-Production
|  | POTENTIALLY | Assembly Correct (Part 3 ) Screws sealed (Screws which require seals (per spec's) are sealed and sealed securely.) | Leak test/ Particle Emitter Test | As part of "Full Assembly" Leak test | Major A | 1 | In-Production OR Post-Production
|  | YES | Cleanliness & Condition: Case + Belt Attachment (Belt attachment loop, Plenum, interior & exterior of Case are free of: cracks, chips, holes, flaking, debris, and any other damage that could compromise integrity of the unit, or cause injury to the user or the user's environment.) | Visual Inspection | N/A | Major A | 1 | In-Production
|  | YES | Company Logo/Product Name is Correct & In Good Condition (Company Logo and Product Name (as applicable) are: present, clear, free of smudging or color bleeding, produced per specs with regard to: color, design, orientation, and size.) | Visual Inspection | N/A | Minor | 4 | In-Production OR Post-Production
10 | BELT | YES | Assembly is Correct (Part 1) Loop, Belt, & Buckle Correct (Belt is attached to blower through attachment loop; female & male parts of belt buckle are oriented per spec's.) | Visual Inspection | N/A | Minor | 4 | Post-Production
|  | YES (if not doing entire exterior Leak Testing) | Assembly is Correct (Part 2) Attachment Loop Fasteners Are to Spec (Screws for belt-to-case attachment loop are securely in place, properly sealed, and tamper-resistant per spec's.) | Visual Inspection | N/A | Minor | 4 | In-Production OR Post-Production
|  | YES (if not doing "Torque testing" OR "Holds Properly With Stress test) | Assembly is Correct (Part 3) Attachment Loop Fasteners Secure (Screws for belt-to-case attachment loop are securely in place and tamper-resistant.) | Torque Limiter/Torque Reader test | N/A | Major A | 1 | In-Production OR Post-Production
|  | YES (if not doing entire exterior Leak Testing) | Assembly is Correct (Part 4) Attachment Loop Fasteners Sealed (Screws for belt-to-case attachment loop are securely in place, properly sealed, and tamper-resistant.) | Part of Full Assembly Leak test/ Particle Emitter Test | Localized Leak test/Particle Emitter test | Minor | 4 | Post-Production
|  | UNSURE (is this more a design verification)? | Belt Loop Secure (Holds properly with stress.) | Stress test (Pull test/ weight test/ Shear test/Torque test at different planes) | N/A | Major B | 2.5 | Post-Production
|  | UNSURE (is this more a Design verification?) | Belt Mobility (Allows belt strap to slide without excessive force.) | Visual / Manual Inspection | N/A | Minor | 4 | In-Production OR Post-Production
|  | YES | Belt Length (Belt is correct length per spec's.) | Supplier CoC | N/A | Minor | 4 | Pre-Production
|  | YES | Belt Strength (Belt is correct length and of standard strength per spec's) | Supplier CoC | N/A | Major B | 2.5 | Pre-Production
|  | YES | Cleanliness & Condition: Belt & Buckle (1) ( Belt & Buckle are free of: debris, frays, rips, tears, flaking, etc.) | Supplier CoC | N/A | Minor | 4 | Pre-Production
|  | YES | Cleanliness & Condition: Belt & Buckle (2) (Buckle's female-male mating pieces connect and remain secure.) | Visual Inspection | Supplier CoC | Major B | 2.5 | Pre-Production
11 | HOOD | YES | Cleanliness & Condition (Part 1): Integrity of Material & Seams (Hood fabric, fasteners, bindings, and seams are clean & free of: rips, tears, flaking, holes (other than the required 10 x manufactured exhalation holes under center of face shield), etc. Face shield is free of: scratches, cracks, dents, chipping, flaking, or visual distortion. Head harness components are free of: rips, tears, holes, cracks, chips, flaking, etc. All seams and bindings are intact, including when all elastic seams are stretched. Any bindings/taping at the seams are smoothly applied, fully adhered, and free of lifted edges or damage.) | Visual Inspection + Manual Inspection | N/A | Major A | 1 | Pre-Production
|  | YES | Assembly Correct (Part 1): Entire Hood Assembly (Hood, face shield, head harness, and hood connector are present and assembled together per spec's. There are 10 manufactured exhalation holes present underneath the center of the face shield. These holes are fully punched out.) | Visual Inspection | N/A | Major A | 1 | Pre-Production
|  | YES | Assembly Correct (Part 2) Head Harness (Head harness parts are per spec's & sewn together correctly. All seams are intact.) | Visual Inspection | N/A | Major A | 1 | Pre-Production
|  | YES | Connectors (Connectors attached to hood are seated correctly and firmly. Hose connects securely & disconnects to hood connector when prompted. No excess force is required to connect or disconnect. Disconnection is not so easy as to occur without intention.) | Visual + Manual | N/A | Major A | 1 | Post-Production
|  | NO (Not needed for PAPR HE) | Communication Test "Determination of Communication Performance Test For Speech Conveyance And Intelligibility." | CVB-APR-STP -0089 | Or equivalent | Minor | 4 | Post-Production
12 | PCB | CONFIRM TYPICAL | Lights are Functional (When PCB connection is tested) (Lights function & remain on or off until prompted not to be, where applicable.) | Visual Inspection | PCB tests | Major B | 2.5 | Post-Production + In-Production?
|  | CONFIRM TYPICAL | Outputs are Aligned to UI (The PCB connections between areas of output and input are correct, to spec, & secured properly.) | PCB tests | ? | Major A | 1 | In-Production
|  | CONFIRM TYPICAL | UI Commands Functional & Responsive (All commands available to the user respond to prompting on the UI board once assembled in the case. The correct responses are produced to the intended commands. The above holds true both: without membrane attached, and after the membrane is securely attached & sealed.) | Visual/Audio Inspection | ? | Major A | 1 | In-Production
|  | CONFIRM TYPICAL | Audible Warnings Responsive to Conditions (Warnings for low battery and low-flow (if applicable) are triggered as input drops at and below specified input levels.) | Visual/Audio Inspection | ? | Major A | 1 | In-Production AND/OR Post-Production
|  | CONFIRM TYPICAL | Command Responses are within Tolerances (Ex. Alarms sound within Db tolerances, alarms deliver correct sound to spec's, shut-on and shut-off responses are within time spec's if specified, etc.) | Decibel testing, Time testing, Visual/Audio Inspection | ? | Major A | 1 | In-Production OR Post-Production
|  | CONFIRM TYPICAL | Soldering Assembly & Condition: (Proper soldering technique is observed and per spec: soldering is intact, in good condition. No soldering missing. No extra soldering where not required. Free of cold joints, dewetting, bridges, etc.) | Visual Inspection | ? | Major A | 1 | In-Production
|  | CONFIRM TYPICAL | Circuit continuity is appropriate | PCB tests | ? | Major A | 1 | In-Production
|  | ANY OTHER TYPICAL ELECTRONICS TESTS | Etc.? | ? | ? | TBD | TBD | TBD
13 | CONTAINER | YES | Package: Cleanliness & Condition (The product packaging container is clean and free of damage. The lid closes securely and remains securely closed.) | Visual Inspection | Other if recommended | Minor | 4 | Post-Production
|  | YES | Package: Contents Correct and Secure (All items required in and on the package/container are correct per specs. The contents are securely packaged to avoid damage to items within and the lid closes securely with all items packed inside.) | Visual Inspection | Weight Inspection? Other if recommended | Minor | 4 | Post-Production

## To Do
Things that need attention (this version)
1. The Battery charger
a. column C, rows 5-6 has to be changed to describe whatever we are using now instead (power source and other cable?)
b. I left a blank space in column E, row 6 under "Functioning" for someone to specify what the functioning will be checked with.
2. Batteries
a. the description for Batteries in column C, rows 7-8 needs to be changed to whatever we're calling the battery pack (or is it just battery pack?)
b. are there other things the Team believes needs to be checked about batteries to ensure quality product is sent out to the user?
3. Airflow test and Noise/Communication Test
a. unsure if we are doing these or not? We still could even if NIOSH doesn't need it for the EUA.
b. we were going to ask Tom if we needed to do the Airflow one as that seems to be for another type of PAPR design than ours.
4. Belt attachment
a. rows 47-48 mention "Tamper-resistant" screws. I don't think we are doing these anymore, right?